## Slide 1: MiniLab Series for Industrial Plants
March 2016

## Slide 2
In service oil analysis for Machine Condition Monitoring
Early identification and trending of machinery wear and failure modes allows maintenance before catastrophic failure occurs.
Oil analysis complements vibration, thermography and other predictive technologies
Increases confidence in decision to remove machinery from service.

## Slide 3
Oil Analysis answers the critical questions ...
Trivector graph of machinery health

## Slide 4: Typical limits for machinery types

## Slide 5: Key Analysis Parameters of the MiniLab 153

## Slide 6
Elemental Analysis
Viscosity
Chemical
Analysis
Particle Count and Ferrous
4 simple tests to quantify machinery health
MiniLab 153

## Slide 7: Four simple tests - Elemental Analysis
  Test 1: Elemental analysis
    Sharpen electrode and install with disk
    Pipette oil sample into cup
    Run sample
Load sample into cup
Sharpen electrode
Insert electrode and disk
Test

## Slide 8: Four simple tests - Particle Count and Ferrous
  Test 2: Particle Count and Wear
    Dilute sample if viscosity >320 cSt
    Shake 30 seconds by hand
    30 seconds in ultrasound to remove bubbles
    3 minutes test through LNF particle analyzer
Ultrasound
(deaeration)
Dilute
Sample
Shake
by hand
Test

## Slide 9: Four simple tests - Viscosity & Chemistry
Test 3: Viscosity
Pipette oil sample,drop 60 uL into viscometer with positive displacement pipette
Not homogenized
Test 4: Chemistry and Water
Homogenize sample, pipette 2 drops onto FluidScan
Homogenization finely disperses free water so it can be accurately measured by the FluidScan

## Slide 10
OilView Alarm analytics
Review Alarm history by individual component or class of components
Adjust initial alarm limits using the OilView cumulative probability distribution analysis capability
Analyze correlations between Analysis Parameters for insight into root cause

## Slide 11: Trivector Reports
Trend Report with automatic import of Wear Image ‘thumbnails”
Optional WDA Module imports images from microscope camera or from LNF export file
-Export all images or filter by class
All traditional OilView
reports still available

## Slide 12: Performance verification with standard fluids
Total Ferrous and Ferrous Particle Count
LNF-545 Ferrous Validation Standard for Magnetometer
  Total Ferrous Content, nominal 100 ppm
  Large Ferrous particle count, nominal 42 um particles
Particle count
LNF-509 calibration check fluid
Fluidscan IR check fluid
FL311 FluidScan Check Fluid
Viscosity standards
  PV1025 Visc 30 standard
  PV1026 Visc120 standard
Elemental
CS-24 standard
NIST traceable standards and an ASTM method for instrumentation

## Slide 13: Minilab series configurations
Three models with increasing capability
Minilab 23
Minilab 153
Minilab 53

## Slide 14: MiniLab parameters
A solution for every oil program with a clear upgrade path

## Slide 15: Excellent repeatability of measurements for trending
 | Analytical Range | Repeatability
Particle Count | 4 um-100um | ± 6% RSD
Total Ferrous | 10 - 2,000 ppm | ± 5% RSD
Ferrous Particle Count | 25-100 um | ± 5% RSD
Viscosity (Low) | 1-320 cSt | ± 3% RSD
Viscosity (High) | 320-700 cSt | ± 5% RSD
TAN | 0-6 mg KOH/g | ± 3% RSD
Oxidation | 0-3 abs/mm2 | ± 3% RSD
Dissolved water | 200 ppm-saturation | ± 6% RSD
Free water | 0.1-6.5% (1,000-65,000 ppm) | ± 25% RSD
Elemental | Varies with element | Varies with element

## Slide 16: Lab report vs. OilView® TriVectorTM report
 | Typical lab report | Trivector report
Elemental analysis | Yes | 
Particle count & ISO codes |  | 
Viscosity,40C |  | 
Water |  | 
TAN/TBN |  | 
Oxidation |  | 
Total Ferrous | PQ/DR ferrography | Fe in ppm
Large Fe count & distribution | DR ferrography | Fe distribution >20um
Analytical ferrography | Optional | Automated classification
Wear images | X | Yes
Wear classification (cutting, sliding, fatigue) | X | Trendable data
Non-metallic count & distribution | X | Trendable data
Simplified for on-site oil analysis - automated ferrography for the non-expert

## Slide 17: How does LNF particle analyzer work

## Slide 18: How does the RDE Elemental analyzer work
Oil sample is loaded into cup. Rotating disk carries oil into AC arc
Emission lines emitted representative of specific elements
Spectrum captured with CCD detectors

## Slide 19: How does Fluidscan® work
IR Spectrum
Measured fluid properties
Built-in Fluid Reference Library

## Slide 20: How does the viscometer work
Elapsed time for sample to travel under gravity down 40C heated cell is measured and the Kinematic Viscosity calculated
Viscosity= A* (t1 - t0) + B
  *A and B are calibration coefficients
Oil Flow by gravity
Time = t0
Time = t1
Temperature = 40C

## Slide 21
21
21
MiniLab 153 analyses
The Infrared Spectrometer measures TAN (Total Acid Number), Water and Oxidation for machinery oils and hydraulics.
For engine oil it measures oxidation, nitration, sulfation, total base number (TBN), soot and water
The Optical Emission Spectrometer measures 23 elements: wear metals from mechanical components, elements in oil additives and contamination.
The Viscometer measures 40C kinematic viscosity.
It also calculates 100C viscosity using the Viscosity Index (VI).
The Particle Counter, Ferrous Monitor and Wear Classifier provides:
Particle Count & ISO Contamination Codes
Total Ferrous content and Ferrous particle count
Wear particle count and classification

## Slide 22
The MiniLab 153 delivers comprehensive on-site oil analysis with:
4 simple tests and results in < 20 minutes
Lab-quality analysis results in your plant
Simplified workflow for the non-expert user
Trivector reports with immediate actionable results
Full integration to AMS OilView software