## Tools Spec
Products | FieldLab Q5800 | MiniLab 153 | Microlab 40
Pictures
Tools consists of | Portable Viscometer Q3000 | Portable Viscometer Q3050 | Dual Temperature Viscometer
| Fluidscan Q1000 IR Spectrometer | Fluidscan Q1100 IR Spectrometer | IR Spectrometer
| XRF Elemental Analyzer | Q100 RDE Elemental Analyzer | OES Elemental Analyzer
| FPQ Particle Counter | LNF Q230 Particle Counter | Light Blockage Particle Counter
Parameter of Analysis | Kinematic Viscosity at 40 degree C (cSt) | Kinematic Viscosity at 40 degree C (cSt) and calculated kinematic Visocisity at 100 degree C (cSt) | Kinematic Viscosity at 40 and 100 degree C (cSt)
| TAN & TBN (mgKOH/g) | TAN & TBN (mgKOH/g) | TBN (mgKOH/g)
| Oxidation, Nitration, Sulfation (Abs/0.1mm) | Oxidation, Nitration, Sulfation (Abs/0.1mm) | Oxidation, Nitration (Abs/cm)
| Water Content (ppm) | Water Content (ppm) | Water, Soot & Glycol (% weight)
| Soot & Glycol (% weight) | Soot & Glycol (% weight)
| Silicon (Si), Aluminum (Al), Chromium (Cr), Titanium (Ti), Iron (Fe), Nickel (Ni), Lead (Pb), Copper (Cu), Tin (Sn), Molybdenum (Mo), Silver (Ag), Zinc (Zn), Vanadium (V) in ppm | Silicon (Si), Aluminum (Al), Chromium (Cr), Titanium (Ti), Iron (Fe), Nickel (Ni), Lead (Pb), Copper (Cu), Tin (Sn), Molybdenum (Mo), Silver (Ag), Zinc (Zn), Vanadium (V), Barium (Ba), Boron (B),Cadmium (Cd), Calcium (Ca), Magnesium (Mg), Manganese (Mn), Phosporus (P), Sodium (Na), Lithium (Li), Potassium (K), Antimony (Sb) in ppm | Basic Metals : Silicon (Si), Aluminum (Al), Chromium (Cr), Iron (Fe), Lead (Pb), Copper (Cu), Tin (Sn), Molybdenum (Mo), Potassium (K), Sodium (Na) in ppm. Extended Metals : Zinc (Zn), Vanadium (V), Barium (Ba), Boron (B), Titanium (Ti), Nickel (Ni), Calcium (Ca), Magnesium (Mg), Manganese (Mn), Phosporus (P) in ppm
| Particles > 4 micron/ml on filtergram | Particles > 4 micron/ml | Particles > 4 micron/ml
Wear Classifier (Cutting, sliding, fatigue, fibers, non-metal)
Total Ferrous (Fe) in ppm
Ferrous particle size & distribution
Price | available on request | available on request | available on request
Advantages | All in one on-field rugged analyzer, suitable for remote area usage where th samples can't travel or very hard to send to Lab. | All in one on-site analyzer with a very detailed results that can be used as root cause analysis data. | All in one on-site analyzer, with data trending and analyzer built into the unit.
| AMS Oil View compatible | AMS Oil View compatible | Automated testing procedure, operators only needed for placing sample on the analyzer and inputing data that are needed for testing
Disadvantages | XRF Elemental Analyzer requires a certification (depends on local goverment's regulations for working with radiation tools) | Most expensive from all of the tools compared in this table. | No Total Acid Number (TAN) results, which is needed for non-engine equipment Oil Analysis. Adding Fluidscan Q1100 which can produce TAN result will take away the advantage of this unit automated testing procedure.

## Tools Technology
Products | FieldLab Q5800 | MiniLab 153 | Microlab 40
Viscosity | ASTM WK40831 - "New Test Method for Field Determination of Kinematic Viscosity Using Microchannel Viscometer" | ASTM WK40831 - "New Test Method for Field Determination of Kinematic Viscosity Using Microchannel Viscometer" | Dual Temprature Viscometer
| Measures Kinematic Viscosity at 40C | Measures Kinematic Viscosity at 40C and calculate Kinematic Viscosity at 100C | Measures Kinematic Viscosity at 40C and 100C
TAN & TBN, Oxidation, Nitration, Sulfation, Water Content, Soot & Glycol | ASTM D7889 Compliant- “Standard Test Method for Field Determination of In-service Fluid Properties Using IR Spectroscopy” | ASTM D7889 Compliant- “Standard Test Method for Field Determination of In-service Fluid Properties Using IR Spectroscopy” | Infrared Spectrometer
| Infrared Spectrometer | Infrared Spectrometer
Elemental Analysis | ASTM WK38220 "New Test Method for Coupled Particulate and Elemental Analysis using XRF for In-Service Oils" | ASTM D6595 "Standard Test Method for Determination of Wear Metals and Contaminants in Used Lubricating Oils or Used Hydraulic Fluids by Rotating Disc Electrode Atomic Emission Spectrometry"
Wear particle size measured around 5-10 microns